## Slide 1: Design challenges of wearable healthcare and patient monitoring
Kelvin Le
SEM – Medical
Notes: Hello and Welcome to this webinar. My name is Kelvin Le, and I’m Systems Engineer here at Texas Instruments. I hope all of you are staying safe and healthy in this difficult time. I would like to thank each one of you for taking out time for this webinar. One request from my side. Let’s make this session interactive.

## Slide 2: Agenda
TI in medical
Market trends
Fundamentals and challenges
Reference designs
Notes: Here is the high-level agenda of this webinar. We will start with an introduction to TI’s classification of Medical sector. Then we will talk about market trends and design challenges, and lastly we will discuss about TI’s subsystem level reference designs which can help in quick prototyping and reducing the design cycles.

## Slide 3
Home healthcare
Medical equipment
Patient monitoring & diagnostics
Medical imaging
TI semiconductors in every medical category
Personal care & fitness
Notes: TI classifies medical market into 5 different categories: Imaging, Patient monitoring and diagnostics, Medical equipment, home healthcare and personal care and fitness.
Medical imaging consists of equipments which are related to some image of the patient. For example, Ultrasound scanners, X-ray machines, MRI and CT/PET scanners.
Patient monitoring and diagnostics category takes care of all the vital sign sensing and monitoring.
If we consider patient monitoring category as a category which takes human vital signs as “Inputs”, then Medical equipments are mostly “Output devices” like ventilators, infusion pumps and anesthesia machines.
The fourth category is home healthcare category and it covers devices which are mostly used at home like blood pressure monitors, glucose monitors, CPAP machines, nebulizers and oxygen concentrators.
And the final category, Personal care & fitness includes devices like wearable fitness monitors, beauty and grooming like shavers etc, and treadmills.
In this presentation, we will talk about the Patient monitoring and diagnostics category.
Click animation

## Slide 4: Medical sector page
Notes: Before diving deeper I would like to show where you can find all the Medical related content on our website.
Once you type www.ti.com in your browser, it will show our home page and one can click on Applications section on top left corner.
It will show different markets and sectors which Texas Instruments supports.
Just click on Medical sector within Industrial market and it will open up our medical portal.
As I highlighted earlier in this webinar, it will show 5 different categories within medical sector.
Each blue color text corresponds to a product or equipment and is clickable.
Let’s take an example. Slide-4
Click animation
Click animation
Click animation

## Slide 5: Medical sensor patches
Notes: Within Patient Monitoring & diagnostics, if I click on “Medical Sensor Patches”, it opens up key design challenges and a high -level block diagram.
Each block here is also clickable with more information – like reference designs for that particular subsystem and recommended products from Texas Instruments.
Slide-5
Click animation
Click animation
Click animation
Click animation
Click animation

## Slide 6: Reference design
Comprehensive
design guides
Detailed design
considerations &
applications info
Full
schematics
Design files
TINA-TI
spice simulation
Bill
of
materials
Comprehensive
test results
Boards
available for
evaluation
Notes: So what is a reference design? A reference design is a subsystem level circuit or design which combines multiple integrated circuits from Texas Instruments and solves a design challenge.
If you are aware of evaluation modules from TI, then reference design is not an evaluation module. It is a resource for our customers with comprehensive design guides, calculations, equations, full schematics, layout, gerbers and design files.
As the name suggests, it is a tested reference to help customers in reducing the design cycle. Remember, It is not a full product.
Slide-6
Click animation

## Slide 7: Patient monitoring market trend
Remote monitoring enhances quality of care and reduces healthcare cost
Wearable wireless medical technology enables accurate and reliable data in a smaller form factor: multi-modalities, longer battery life, SHIP mode
Artificial Intelligence uses analytics and big data to improve decision making and early prevention
Notes: Patient monitoring is not new but the market is growing.
For example, the growing elderly population drives the high demand of remote monitoring.
In many countries, Billions of dollars are being spent every year on chronic diseases.
Remote monitoring has become trend and it not only enhances the quality of care but also reduces overall healthcare cost while providing mobility advantage to the patient.
Wearable wireless medical technology enables accurate and reliable data in a small form factor.
The challenges to tackle with wearable medical devices are: the battery life, the variety of vital signs or modalities in smaller form factor design.
Once the vital signs are monitored, captured and sent for processing, Artificial intelligence can be used for analysis. It can improve the decision making and help in early prevention of the diseases. Slide-7

## Slide 8: Patient monitoring basics
The electrocardiogram (ECG)
measures electrical activity of the heart
Photoplethysmography (PPG) is an optical measurement of an organ’s volume.
Notes: Before we go into details, let’s briefly talk about ECG, PPG and some of physiologies behind it.
ECG is stands for “electrocardiogram,” which is an electrical measurement of the heart activity.
During each heartbeat, the cardiac muscle tissue contracts in a specific sequence in order for blood to flow in the proper direction, passing from one chamber to the next.
The contraction of each segment in the heart produces its own depolarization waveform, like the ones shown in the figure on left.
PPG stands for “photoplethysmogram,” which is an optical measurement of an organ’s volume.
In principle, PPG is measured by illuminating the skin and subcuntaneous tissue with light of a specific wavelength from an LED or light-emitting diode.
This light is either absorbed, passed through, or reflected back.
A photodiode sensor on the other side measures the light that is either transmitted or reflected, depending on where it is placed relative to the LED.
The light is then converted to an electrical signal.
Pulse oximeter is the most common device used to measure PPG. These devices use two specific wavelengths of light: Red (660nm) and Infrared (940nm). The absorption of light at each wavelength allows physicians to isolate the light absorbed by the arterial blood from venous blood and other types of tissue. This is primarily used a an indicator of the oxygen concentration in a patient’s bloodstream.
When measured simultaneously, we can overlay ECG and PPG waveforms to see how the two signals are related.
Notice how both signals have the same period, meaning that either one can be studied to determine heart rate.
The time difference between the R-peak in the ECG waveform and peak of PPG waveform is known as Pulse Transit Time or PTT.
PTT involves simultaneous measurement of ECG and PPG together.
Along with other variables, such as the patient’s size, weight, age, etc., researchers have been writing algorithms that show the correlation between PTT and systolic blood pressure.
Slide-8

## Slide 9: ECG vs. PPG
Feature Description | ECG | PPG
Measurement type | Electrical | Optical
Sensor type | Electrodes | Photodiode
Can measure heart rate? | Yes | Yes
Diagnostic information | Yes | Yes
Minimum number of skin contacts required? | 2 (Across chest) | 1 (Finger or wrist)
Number of ADC channels required | ≥1 | 1
Notes: This table summarizes the differences between ECG and PPG measurements. It compares features of both ECG & PPG measurements. I am not going to talk about each features, but it is important to highlight that ECG & PPG have their own benefits and ways to provide diagnostics information.
Slide-9

## Slide 10: ECG lead and ADC channels
Standards Electrodes Needed
1 Lead LA, RA
3 Lead LA, RA, LL
6 Leads LA, RA, LL
12 Leads LA, RA, LL, V1-6
Number of Leads | Leads Used | Number of ADC Channels
1 | Lead I | 1
3 | Lead I, Lead II, Lead III | 2
6 | Lead I, Lead II, Lead III, aVR, aVL, aVF | 2
12 | Lead I, Lead II, Lead III, aVR, aVL, aVF, V1 – V6 | 8
Notes: A “lead” in ECG is the voltage difference between two points on the body.
For most leads, each point is measured by an independent electrode, though some can be calculated.
The number of leads in an ECG system can vary widely.
Simpler ECG systems generally use no more than the first two or three primary leads. This is typical in consumer applications and small, portable ECG equipment meant for basic diagnostic purposes.
On the other hand, high-quality diagnostic systems are more complex and may use up to 15 different leads!
This table summarizes what leads you measure for single-lead, three-lead, six-lead, and 12-lead ECG, which are the most common configurations. Also shown is how many ADC channels are required to measure each one. Slide-10

## Slide 11: ECG characteristics Frequency domain
Notes: The ac component of the ECG waveform is relatively low in frequency, usually between 0.05Hz and 40Hz.
Diagnostic quality ECG applications may require up to 150Hz or more to extract additional information from the waveform.
This diagram shows where ECG falls in the frequency spectrum relative to other biopotential measurements.
Again - note the magnitude of ECG is just a few millivolts.
Slide-11

## Slide 12: Challenges in measuring ECG
Notes: As we alluded to earlier, there are some common challenges to measuring ECG. These challenges usually come from power line interference or also called 50/60 Hz pick-up, baseline dc instability which comes from poor electrode connection, muscle shaking, and wandering baseline from changing dc level of the electrodes.
Some of these challenges can be overcome with a good circuit design, but most often, digital post processing is required to bandpass filter the ECG waveform from DC drift or high-frequency interference.
Slide-12
Click animation
Click animation
Click animation

## Slide 13: Challenges in optical bio-sensing
Low power for longer battery life
Skin tone variation
Best PPG signal for motion cancellation algorithms
Performance with glass
Low temperature performance
Ambient light
Notes: So what are the challenges for optical bio-sensing Apart from longer battery life, which is a common challenge with all the wearable products?
As highlighted earlier, PPG measurement uses LED transmitting the light signal and photodiode receiving the reflected signal.
The patient’s skin tone variation can really affect the receieved signal. Its amplitude is highly dependent on skin color.
While combining the LED & photodiode circuit in a small form factor, the performance with glass and lense is important along with cancellation of ambient light and correction in software.
Many wearable monitors also have motion cancellation algorithms to compensate for any motion artifacts.
So let’s dive deeper into the solutions from Texas Instruments which can help solving these challenges.
With this background, I would like to talk about few reference designs.
Slide-13

## Slide 14: TIDA-01614 Multiparameter front-end for vital signs patient monitor reference design
Benefits
Single IC does both ECG ,Respiration.
Pace Detection
ECG with 3 electrodes
Three temperature sensors for temperature measurement
Features
System measures ECG, Heart Rate, SPO2, Respiration rate using ADS1292R and AFE4403 and Skin temperature using TMP117
Circuit enables three electrode operation including right leg drive with good CMRR
Pace detection circuit indicates presence of pacemaker
Supports three 0.1 Celsius accurate sensors (TMP117) to measure the skin temperature
Enables data transfer over isolated UART interface
Works with 3.7V Li ion rechargeable battery
On board memory for data logging
Applications
Multiparameter Patient Monitor
Medical Sensor Patches
Pulse Oximeter
Electrocardiogram (ECG)
Notes: Let me start with our first reference design TIDA-01614. It is a reference design for multiparameter front-end for vital signs patient monitors.
It measures ECG, heart rate, SPO2, respiration rate and skin temperature.
Since the block diagram on this slide is very small and not clearly visible, I will blow it up and show in details. Slide-15

## Slide 15: Detailed block diagram for TIDA-01614
15
Notes: The ECG, heart rate, and respiration rate – these three parameters are measured using dedicated ECG chip ADS1292R.
The circuit enables three electrode operation including right leg drive with good common-mode rejection ratio.
The SPO2 is measured using dedicated PPG IC AFE4403. It uses Transmissive type of PPG measurement and can take input from a finger-clip connector.
For measuring skin temperature, the reference design supports three temperature sensors through external board connection.
The temperature sensor boards use TI’s latest temp sensor TMP117 which is +/-0.1 degree accurate sensor.
There is also an option to connect pace detection circuit to indicate presence of pacemaker.
This reference design operates either from a 5V USB input or from a single cell Lithium Ion battery and has battery charging and other power components onboard.
The raw data for all vital signs are available over isolated UART interface shown on the right side.

## Slide 16: Design challenges TIDA-01614 solves
Design challenge 2
Protection and isolation
Design challenge 1
Integration of multiple modalities at optimum SNR levels and small form factor
Monitoring of ECG, Heart Rate, SPO2, PTT, Respiration rate and Skin temperature
Single Lead ECG with RLD (ADS1292R)
> Signal amplitude: 0.2mV~2mV (p-p);
> BW0.05 Hz to 2000 Hz
Supports 3 LED and 3 Photodiodes with ambient subtraction for SPO2 and Heart Rate monitoring with AFE4403
Supports three 0.1 Celsius accurate sensors to measure the skin temperature (TMP117)
Isolated UART interface using an onboard MSP432P401, ISOW7842, TRS3232
Notes: What design challenges does this reference design solve?
It actually helps in two important aspects. One is integration of multiple modalities with enough bandwidth, signal amplitudes and signal to noise ratios.
Second one is protection and isolation to support patient safety.
Slide-16

## Slide 17: ECG analog front end
  TIDA-01614: Multiparameter front-end reference design for vital signs patient monitor
Important parameters:
Input bias current
Input impedance
Input current noise
Input voltage noise
Power consumption
DC/AC CMRR
Notes: To highlight more on protection, the signal chain circuit is protected using resistors for limiting the current on the electrodes , common-mode and anti-aliasing filters to reject high-frequency noise, and diode circuit to protect against defibrillator shocks.
Also it is important to understand that DC specifications like offset, gain error, and nonlinearity do not matter because the offset from the skin-electrode contacts are typically much larger than errors introduced by an instrumentation amplifier.
However, input bias current, impedance and noise numbers are important.
The right leg drive amplifier has two purposes – it biases the patient to the proper voltage range and cancels out common-mode interference. The reference design board is of a cell phone size which can be worn on arm or on waist while supporting patient mobility.
Slide - 17

## Slide 18: TIDA-01614 test setup and test results
Test setup
GUI display
Design specs
Notes: This slide shows the design specifications, test set-up showing how it is tested using ECG simulator and a graphical user interface to show all different parameters.
Just remember that GUI, schematic, layout, gerbers and test results – everything is available on TI web. You need to just search for TIDA-01614.
Slide-18

## Slide 19
Test pacemaker detection with TIDA-010005 & TIDA-01614
Notes: Here the same GUI shows ECG, respiration, SPO2 and three temperature signals. In fact, it also shows an indication of pace pulse along with rise time, amplitude, and width.
Now let’s move to wearable patient monitor reference designs.
Slide-19
Click animation

## Slide 20: TIDA-01580 Wearable, wireless, multiparameter patient monitor reference design
Device Datasheets:
  AFE4900
  CC2640R2F
  TPS61098
  TPS63036
  TPD1E10B06
PPG supports 4 LEDs and 3 PDs with Digital Ambient subtraction to improve the SNR
AC and DC lead off detection helps in correct measurement of vital signs
Continuous Monitoring with lower operating power ensures battery life of 30 days
Flexibility of ultra low power modes and integrated FIFO can keep MCU into sleep to increase the battery operation time
Simple Wearable Multi-Parameter Patient Monitor for Photo-plethysmography (PPG) and Electrocardiography (ECG)
Provides Raw data to calculate heart-rate, Oxygen Concentration in Blood (SpO2) and Pulse-transit Time (PTT)
Uses Single-chip Bio-sensing Front-End AFE4900 for Synchronized ECG & PPG
  PPG (Optical heart-rate monitoring and SpO2) supports 4 LEDs and 3 PDs with Digital Ambient subtraction to improve the SNR
  ECG (LEAD I) signals
Integrated ARM Cortex-M3 + 2.4GHz RF Transceiver (CC2640R2F) supports wireless data transfer – BLE 4.2 and 5
Operated from CR3032 (3V, 500mA Coin Cell Battery) with battery life of 30 days using highly efficient DC/DC converters
Small form factor helps in easy adaptation to wearable applications
Features
Benefits
Wireless patient monitor
Pulse Oximeter
Target applications
Tools & resources
Wearable fitness & activity Monitor
ECG
Notes: The TIDA-01580 reference design realizes a wearable patient monitor.
It is very simple, multi-parameter, patient monitor that provides raw data to calculate heartrate, SpO2, and Pulse transit time.
The Single-chip, biosensing front-end AFE4900 supports synchronized ECG and PPG measurements.
It is the only device on the market with this level of integration, low power operation, small size, and flexibility.
The PPG or the optical heart-rate monitoring and SpO2 supports 4 LEDs and 3 Photo diodes with digital ambient subtraction to improve the Signal to noise ratio.
Along with PPG, it supports single lead ECG.
The Wireless data transfer over Bluetooth low energy 4.2 and 5.0 are supported by CC2640R2F.
It is an Integrated Arm Cortex-M3 device with 2.4-GHz RF transceiver.
It has internal DC/DC converter which helps improve the overall efficiency, built-in low-battery detection algorithm helps reduce the external components and all digital peripheral pins can be routed to any GPIO which helps in better routing in small form factor.
The AFE4900 has flexibility of ultra-low-power modes and integrated FIFO which can keep the MCU in sleep mode to increase the battery operation time.
The reference design is operated from CR3032(3-V,500-mA coin-cell battery),which offers 100 hours of continuous operation.
If the operation is duty cycled for 1 second/minute transmission then the battery life can go up to 30 days.
Let me highlight different boards on his design.
It has three boards, sensor board, ECG board and main board.
Sensor board has LED & PPG sensor, ECG board has just metal contacts and both these get connected to main board.
It can support multiple PPG sensors because the PPG board is plug-in type.
The design has been tested with PPG sensors from different vendors – one is 3 LED sensor and other one is 4 LED sensor.
As shown in left down corner, there is one metal contact on top layer and another on bottom for two electrode ECG measurement. There is a small window from where PPG sensor measures does the optical measurement.
Slide-20

## Slide 21: TIDA-01580 for medical patch
Bottom side is touching the wrist of one hand (ELECTRODE 1)
Other hand is touching the PAD on the top layer of the main board.
(ELECTRODE 2)
Side View
LAUNCHXL-CC2640R2F receives the signals remotely and displays on LabView GUI
The design uses BLE 5.0 with an advertising time = 100ms
Notes: This is how the test set-up looks for TIDA-01580. Bottom side can touch the wrist of one hand (let’s say ELECTRODE 1). And Other hand can touch the PAD on the top layer of the main board.(which is ELECTRODE 2).
The raw data is captured and send to remote terminal which can show synchronized PPG & ECG data. Here we have used Launchpad for CC2640R2F as the receive terminal.
Slide-21

## Slide 22: Design challenges TIDA-01580 solves
Design challenge #1
Integration of multiple modalities at optimum SNR levels and small form factor
Capturing synchronized ECG and PPG to enable PTT and BP calculations (non-invasive and without cuff) Pulse Transit Time (PTT): Time difference between the R-peak in the ECG waveform and the arrival of the blood pressure wave Simultaneous measurement of ECG and PPG together Along with other variables, such as the patient’s size, weight, age, etc., algorithms show the correlation between PTT and systolic blood pressure. Challenging to synchronize both measurements – timing is the key! (Powering up, clock timing, phase, drift with temperature) | 
22
Notes: So what design challenges does this design solve?
First and foremost the integration of modalities in a small form factor.
Capturing synchronized ECG and PPG to enable PTT and BP calculations (non-invasive and without cuff) can be challenging. It is mainly because timing is the key! There are multiple aspects involved like powering up, clock timing, phase and drift with temperature. The
Simultaneous measurement of ECG and PPG together along with other variables, such as the patient’s size, weight, age, etc., can be used for further analysis.
Slide-22

## Slide 23: Design challenges TIDA-01580 solves
Design challenge # 2
BLE connectivity that does not interfere with measurement accuracy
Signal amplitude: 0.2mV to 2mV (p-p) BW: as broad as 0.05 Hz to 300 Hz (Pace detection increases the bandwidth further) Reject environmental electrical signals, such as ac mains, security systems, and RFI to amplify and display the ECG signal Good CMRR of the signal chain and Right-leg drive (RLD) for CM rejection Differential- and common-mode filtering, environmental shielding, and algorithms | 
Notes: The second challenge is RF intereference due to Bluetooth connectivity.
The signal amplitude are very low typically 0.2mV to 2mV (p-p).
It is important to reject environmental electrical signals, such as ac mains, security systems, and other RF interference and to amplify and display the ECG signal. This reference design implements differential- and common-mode filtering using common-mode chokes. Also, the board is 4-layers with a complete ground layer for shielding. The placements of BLE chip and AFE chip are done such as the interefence is not there.
Slide - 23

## Slide 24: Design challenges TIDA-01580 solves
Design challenge # 3
Extending battery life to multiple days to enable portability & wearability
Selecting the right power architecture to enable extended battery life up to 24 hours (for rechargeable batteries) or 7 days (for primary cells) Powering with right buck, boost or buck-boost device instead of directly powering from battery (bypass modes in DC/DC converters) Sleep / shutdown / standby / deep sleep modes for radio devices like BLE, Wi-Fi etc. Selection of right battery charger (charging rate, termination current and quiescent currents are important!) | 
http://www.ti.com/lit/an/slyt763/slyt763.pdf
Notes: The third and very important design challenge is to improve the battery life. This can be done in multiple ways, for example selecting the right power architecture, powering with right buck, boost or buck-boost devices instead of directly powering from battery. There are dc/dc converters available which have bypass modes to improve efficiency.
The Sleep / shutdown / standby / deep sleep modes for radio devices like BLE, Wi-Fi etc. can also help improve the battery life. Selection of right battery charger like charging rate, termination current and quiescent currents are important to enahce battery life further.
Here is a journal article which was published last year. It describes all these different techniques to improve the battery life in wearable patient monitors.
So we talked about PPG, ECG & PTT measurement in this patch design. Slide-24
Click animation

## Slide 25: TIDA-01624 Bluetooth-enabled high accuracy skin temperature measurement flex PCB patch
TIDA-01624 and/or Tools Folder
Design Guide
Design Files: Schematics, BOM, Gerbers, Software, etc.
Device datasheets:
  TMP117
  CC2640R2F
Low power consumption and long battery life
Extremely long shelf life (3+ Years)
Small, Flexible Form Factor
Connects to Smart Device
Zero-Calibration to ±0.1°C Accuracy
High Accuracy, Low Power Temperature Sensor
BLE 4.2 and 5 enabled microcontroller
Thin-Film Flexible Battery Power, enabling entirely flexible design
Integrated PCB antenna
Temperature updates every second
Features
Benefits
Tools & resources
Medical Sensor Patches
Multiparameter Patient Monitors
Smart Patches
Applications
Skin Temperature TMP117
BLE MCU CC2640R2F
I2C
Notes: What about temperature patch? so here it is. TIDA-01624.
It is Bluetooth-enabled high accuracy skin temperature measurement flex PCB patch. It is a simplest patch design with only two ICs on it - A temp sensor and a BLE chip.
The temp sensor is again TMP117 which is +/-0.1 degree accurate. The patch is implemented on a flexible PCB that means it can bend and take shape of the body part where is it connected. The silve color battery seen on the picture is also a flexible battery.
The Temperature updates every second and connects to Smart Device like an iPad or iPhone or any remote terminal. More details can be found on TI web by searching TIDA-01624.
Slide-25

## Slide 26: TMP117x Ultra-high accuracy digital temp sensor with integrated non-volatile memory
Ultra-high accuracy
Meets ASTM E1112 & ISO medical standards:
    0.1°C acc. range 35.8°C to 42°C
No calibration needed; NIST Traceable
Integrated non-volatile memory
Store configuration even after losing power
64 Bits of general-purpose scratch pad memory
Low power consumption
  3.5uA Average Iq @ 1Hz; serial bus inactive
  150nA Shutdown Iq; serial bus inactive
Digital feature & I2C interface
Programmable Temperature Alert & Offset value
Soft Device Rest
Smallest package
6 PIN, QFN & CSP
Features
Benefits
Accuracy
  16-bit Resolution (0.0078°C)
  Minimum PSRR: 1LSB = 7.8 m°C/V
Integrated EEPROM
Low power consumption
  140uA Iq during conversion
  3.5uA Average Iq @ 1Hz
  150nA Shutdown Iq
  1.8V – 5.5V
Digital feature: Automatic offset NVM/ Soft Reset
Interface: Single wire
Packaging
  6pin WSON (2 x 2) mm
  6pin WCSP (1.6 x 1) mm
Gas Meter
Medical
Cold Chain
Wearables
Instrumentation & Test
Thermocouple – Reference
Applications
TI Part | Accuracy (°C) | Accuracy Full Range
TMP117M | ±0.1oC @ (30oC to 45oC) | ±0.2oC @ (0oC to 85oC)
TMP117 | ±0.1oC @ (-20oC to 50oC) | ±0.3oC @ (-55oC to 150oC)
TMP117N | ±0.2oC @ (-40oC to 100oC) | ±0.3oC @ (-55oC to 150oC)
Notes: Before we move to next reference design, let me quickly flash this one pager for TMP117.
As it can be seen, it is an ultra-high accurate device and meets medical standards. It has integrated non-volatile memory which can store configuration even after losing power. It is so small with 1.6 mm square packge that it is ideal for wearable application.
Slide-26

## Slide 27: Full system: Multiparameter patient monitor + wireless sensors
Patch #1
Patch #2
Patch #3
Patch #4
Nurse’s station / Doctor’s office
Notes: If we look at overall system, we have multiparameter patient monitor which has connection to nurse station or doctor’s office through WiFi. The sensor patches connected to different patients can communicate with the same patient monitor through Bluetooth.
Here is why you should consider Texas Instruments for wireless connectivity.
The SimpleLink solutions from Texas Instruments are low power. They are easy to use because of multi-role support, they are secure with FIPS 140-level 1 validation and smaller in size because of BAW technology. It is industry’s first crystal-less wireless architecture which saves 12% area in patch designs.
Slide-27
Slide-27

## Slide 28: Why TI SimpleLink™ for multiparameter patient monitor + sensor patch?
  BLE SoC with integrated Ultra low Power Sensor Controller
  Wi-Fi low power IoT
  Best-in-class standby current
  CC3135/CC3235 Wi-Fi modules
  5GHz Wi-Fi to reliably connect to hospital network
  BLE multi-role support, up to 32 simultaneous connections
  FIPS 140-Level 1 validation
  Offload CPU bandwidth – HW crypto accelerators
  Secure boot
  BAW: First crystal-less wireless BLE SoC – 12% area savings in reference design
  Tiny BLE SoC: CC2640R2F – 2.7mm x 2.7mm DSBGA
Secure
Small size
Ease of use
Low power
Notes: On the firware side, you can reuse 100% of the code. All different technologies from Texas Instruments like Bluetooth, Sub-1Ghz, Wifi, Zigbee and multi-standard all can have common software using our SDK. So invest once and reuse effortlessly.
Slide-28
Click animation

## Slide 29: Invest once, reuse effortlessly
Learn more about SimpleLink code portability
SimpleLink Medical Resources
CC2640R2F: How do I design an accurate and thermally efficient wearable temperature monitoring system?
100% code reuse
Common software

## Slide 30: Achieving isolation and enabling patient safety
Notes: Let’s change gears and talk about how to achieve isolation and help ensure safety with power and data architectures in Multiparameter patient monitors.
Slide-30

## Slide 31: Patient safety
Patient safety is a global health priority. Recalling resolution WHA55.18 (2002), which urged Member States to “pay the closest possible attention to the problem of patient safety and to establish and strengthen science-based systems, necessary for improving patients’ safety and the quality of health care”, the seventy-second World Health Assembly (WHA72), in May 2019, adopted WHA72.6, a resolution on ‘Global action on patient safety’. (Source: https://www.who.int/patientsafety/en/)
ECG module
SpO2 module
CO2 module
BP module
Monitor
Notes: By definition, a multiparameter patient monitor measures multiple parameters like electrocardiogram or ECG, blood oxygen concentration or SpO2, temperature, blood pressure, carbon dioxide or CO2 and multiple other gases. These parameters are typically measured using small isolated modules which can be connected to the main unit, called ‘Monitor’. So what you see here is an example image of patient monitor with different connectors where all these different modules get connected.
These type of patient monitors are bedside or classic patient monitors which we see in hospitals and clinics.
There is also trend of moving towards portable patient monitors which are small units which can be worn on waist or arm of the patient. The advantage, of such portable type patient monitors, is patient mobility. However, the challenges for these patient monitors are not only size but also quiescent current and efficiency with thermal performance.
Now Patient safety is a global health priority which is also highlighted by World Health Organization so as to provide enough attention during the design and development of multiparameter patient monitors.
Slide-31

## Slide 32: Isolation requirements and safety limits
IEC60601-1: International basic safety and essential performance standard for electrical medical equipment and medical electrical systems.
  Regional compliance
  Editions and versions
Levels of isolation – patient focus
Spacing – creepage & clearances
Safety insulation for transformers
Leakage current limits
  Isolation at the sensing side
  Isolation at the data/power side
Notes: If you are working on any medical equipment, you would be aware of medical standard IEC60601-1. All medical devices must comply with the basic safety and essential performance requirements defined by this standard.
This standard provides general requirements, in a series of editions and versions and it has evolved over years, through to the 4th edition EMC standard that establishes new “environments” when considering electromagnetic compatibility (EMC) between medical electrical equipment and other electronic devices.
With the United States, Canada, and Europe now having fully adopted the 4th edition EMC requirements, the patient monitor manufacturers intending to ship products to these regions must ensure their products comply to the latest standard. However, the same is equally true for manufacturers in regions that have not transitioned to the latest editions of the standard. While transition dates may not be clear in various territories around the world, there is still incentive for manufacturers to comply to the global standards to avoid being caught out by different regional compliance.
The IEC 60601 standard has a long history with a number of revisions. Since 1977 till date, there have been multiple editions and versions adding more protection features as well as safety of the patient.
In the context of safe operation of medical equipment, the 2nd edition of IEC 60601-1 established risk guidelines that applied when a device is within a 6-foot radius from the patient, or called “patient vicinity.”
And three use categories are defined:
Type B (body) equipment operates within this vicinity, but without patient contact. Examples of type B equipment include x-ray machines, and MRI scanners.
Type BF (body floating) equipment makes physical contact with the patient. Examples of type BF equipment include blood pressure monitors, ultrasound equipment, and thermometers.
Type CF (cardiac floating) makes physical contact with the heart. Examples of type CF equipment include defibrillators and dialysis machines.
These types body, body floating or cardiac floating decide applicable levels of isolation, insulation, creepage, clearance.
Specific tables are given in the IEC standard for the creepage and clearance requirement based on corresponding working voltages across insulation or upper rated input voltage which ever is higher. The clearance calculation needs consideration of a multiplication factor if the altitude is greater than 3000 meters.
Also, power supplies use transformers for isolation purpose. So for transformer construction, a reinforced or basic isolation requirement is determined based on which type of insulation is used. For example, single or double or triple layer of insulation sheet will decide the thickness as well as isolation level.
Most of the AC/DC converters are isolated, then why do we need isolation for DC/DC power supply?
By providing up to 5000 V isolation, a DC/DC converter can still protect the patient in the event of a failure of the main AC/DC power supply, thereby avoiding mains voltage levels appearing at any patient points.
The sensing side isolation is typically taken care by reducing the leakage current by adding resistors in the ECG sensing paths for example.
The purpose of data & power side isolation is two-fold. For modules like ECG & temp, there has to be isolation from patient to the circuit to make sure the leakage current passing through the patient’s body is within safe limits. Secondly, when it connects to the main monitor, we need to understand that multiple other modules are also connected together so providing isolation in each module is important.
Slide-32

## Slide 33: Data and power isolation
Notes: To get into further details, let’s have a look at an example module. The image on this slide shows a generic block diagram of ECG module. It highlights the isolation between the analog circuit and the digital processing unit. As highlighted earlier, Medical standard IEC60601-1 specify that multiparameter patient monitors need to have isolation up to 5kV to separate the grounds, prevent any ground loop currents, and prevent leakage currents flowing from the patient’s body.
While isolating data, it also needs the same level of isolation as on the power supply. The data isolation can be done using simple digital isolators but the isolated power topology needs a thoughtful analysis.
For power architectures, there are key specifications which help in deciding the topology. Now let me discuss the key challenges.
Slide-33

## Slide 34: Key design challenges
Input voltage ranging from 3.3V to 24V
  Regulated input vs. non-regulated input
Output voltage ranging from 3.3V to 6V
Output power up to 5W
Open-loop or closed-loop (voltage/current)
  achieving < 1% load regulation
Isolation ~1kV to 5kV
Emission (CISPR22/25, IEC60601-1)
Small form factor (New trend – electronics in cable and portable MPMs)
  reduced BoM
Low cost
One size doesn’t fit all the requirements
Possible Architectures
Flyback
Push-pull
Isolated power module
Isolated power and data module
Notes: As we discuss with different medical customers about the isolated power supplies, we get requirements like regulated or unregulated input voltage, output voltage range depending on what is it powering, power levels they would want to achieve, the feedback mechanism – open-loop or closed-loop or in other words output voltage regulation percentage, isolation levels, emission standards and bill of material, form factor and cost also play a major role in topology selection.
One size doesn’t fit all these requirements so we need to dig deeper into possible architectures or topologies.
Slide-34

## Slide 35
PSR flyback topology
Tertiary feedback winding needed
Optocoupler needed
TL431
Parameter | Value
Input voltage (Vin) | 4.5V to 65V (70V max)
Output voltage (Vout) | Adjustable
Output power (Pout) | 7W max
Isolation level | 5kV (can be tuned as per transformer design)
Size | 45mm x 25mm x 11mm (Depends on transformer design)
Output regulation | 1% achievable
Suggested TI devices:
LM5180
LM25180
Conventional Flyback
Simplified transformer design
No optocouplers required
Smaller size
Total BoM = 21 components (including passives)
Tight load regulation achieved
Primary-side regulated flyback
Refer to “Design Calculator” for complete schematics, BoM and simulation results
Notes: Let’s start with Flyback topology. A conventional Flyback converter is a buck-boost converter and it can easily support output power levels up to 10s of watts (depending upon the transformer design). To get the isolated output information back to the Flyback controller (aka PWM controller) in order to accurately regulate the output, the output is fed back using either an opto-coupler based circuit or a tertiary winding. In such cases, the component count and the solution size is fairly large. In fact, reliability can be a concern with opto-coupler based feedback and an additional winding adds to the cost of the transformer.
There is a variation of Flyback called the PSR Flyback topology which eliminates the need of an opto-coupler or tertiary winding and provides extremely tight load regulation using the feature of the resonant ring caused by the magnetizing inductance and the switch node capacitance. Compared to conventional Flyback, it operates in Discontinuous Conduction Mode (DCM) or Boundary conduction Mode (BCM) depending on the load. It, in turn, eliminates any errors from the transformer DC resistance (DCR) or secondary side diode to provide tight regulation.
However, higher primary and secondary-side RMS currents in PSR Flyback result in conduction losses in the magnetic and semiconductors.
TI suggests two new devices which can support this topology – LM5180 and LM25180 (the difference between the two devices is the input voltage range).
The table on the right shows specifications for LM5180 device and it can support up to 7W and can achieve 1% output voltage regulation.
Slide-35

## Slide 36: Design calculator for LM25180
Notes: Now , I want to quickly show a design calculator for LM25180 where the user can design and select component by inputting different specifications and parameters. This one is available on our website for anyone to download and use.
Slide-36

## Slide 37
Push-pull topology
Tight output regulation due to feedback
Total BoM = 46 components (including passives)
Optocoupler based design – reliability
Push-pull topology (Closed-loop)
Push-pull topology (Open-loop)
PMP design: PMP40588
5-10% regulation
No opto-couplers required
Smaller size, total BoM = 10 components (including passives)
Needs regulated input
Parameter | Value
Input voltage (Vin) | 2.2V to 5.5V
Output voltage (Vout) | 5V unregulated
Output power (Pout) | 5W max
Isolation level | 5kV (can be tuned as per transformer design)
Size | 30mm x 25mm x 6mm (Depends on transformer design)
Output regulation | 5 to 10%
Suggested TI devices:
SN6505A
SN6505B
Notes: There is another topology which is very popular for small bias supplies and auxiliary supplies. It is push-pull topology. This topology utilizes the transformer core magnetizing current more efficiently which results in a smaller magnetic solution compared to the conventional Flyback.
For instance, SN6505 is a push-pull transformer driver specifically designed for small form factor isolated power supplies.
It is an open-loop push-pull topology which simplifies the design by eliminating the feedback loop. The total component count is also very less – close to six excluding the output LDO in the first figure on top of the slide.
The transformer driver without any feedback cannot support a wide input voltage range and instead requires a tightly regulated input voltage.
Being an open-loop topology, there is always a concern on output regulation. That is why the typical load regulation for SN6505 is 5 to 10% and that is why an LDO may be used for regulated output or it can create output ringing for no load conditions.
For tighter output regulation, closed-loop push-pull topology can also be used. Similar to a conventional Flyback, this needs an isolated feedback using opto-coupler. Though the power level is ~5W here, the overall component count and size is huge compared to open-loop push-pull topology using SN6505.

## Slide 38
Isolated power module
Vin
Vout
Suggested TI device:
UCC12050
Same (apple-to-apple) EVM configuration: no ferrite beads , no LDO, no stitch capacitors, on 2 layer PCB
Tested to CISPR32 Limit, in 10m chamber, on same day, in same certified lab.
UCC12050 schematic
Scalability
Parameter | Value
Input voltage (Vin) | 4.5V to 5.5V
Output voltage (Vout) | Regulated 3.3V or 5V
Output power (Pout) | 0.5 W
Isolation level | 5kV RMS reinforced
Size | 10.3mm x 7.5mm x 2.65mm
Output regulation | 1.5%
Notes: Now I would like to talk about integrated solutions.
Innovations in transformer design and higher frequency topologies have enabled IC designers to integrate a transformer and silicon into one IC. For the end user you get a small, lightweight isolated DC/DC power supply without having to design a transformer or compromise on system performance.
The traditional approaches like Flyback or push-pull provide a lot of flexibility but there is drawback with discrete approaches. It is that the complete solution – the transformer and other components – occupies a lot of space on the board. Also, designing a stable and efficient isolated power supply can be challenging.
Various latest solutions, integrate a DC/DC converter with a micro-transformer, along with signal-isolation channels, in a single package.
Here is an example of TI’s latest device UCC12050. It is an isolated power module, and it integrates a transformer and DC/DC controller with a proprietary architecture to achieve high efficiency with very low emissions. The Z-dimension with this device is just 2.65mm.
It provides 500mW (typical) of isolated output power at peak efficiency of 60%. If you need more output power, multiple such modules can be used in the application.
The UCC12050 brings many benefits to the wide variety of isolated DC/DC supply applications. It is designed with an EMI-optimized transformer with only 3.5 pF of primary-to-secondary capacitance and a very quiet control scheme. On its own the solution can pass CISPR32 Class B on a two-layer PCB without ferrite beads or LDOs (and we will talk about it in details in an upcoming slide). The device is robust, rated for reinforced isolation of 5 kVrms and 1.2 kVrms working voltage and will operate at 125°C ambient temperature. The family of devices also includes UCC12040, which is rated for basic isolation of 3 kVrms and 800 Vrms working voltage. UCC12050 is targeted for 5-V input, 3.3-V to 5.4-V output applications requiring 500 mW. Applications requiring higher input or output voltages will need to provide pre-or-post conversion.
Slide-38

## Slide 39
Integrated transformer technology benefits
Discrete Solution
38mm
21mm
7.5mm
Total bias area
273 mm^2
53%
Less Area
Total bias area
127 mm^2
UCC12050
3.5mm
Single chip solution (UCC12050) advantages:
Smaller size and low profile
Very low isolation capacitance Cps for better CMTI and less noise
Simplify design with less components and easy board layout
Digital isolator
Isolated bias UCC12050
Notes: Now let’s compare the size. The total area while using UCC12050 is 53% less than discrete solution. The total area is just 127 mm^2 compared to 273 mm^2 solution. Apart from the size, it simplifies the layout because of very less external components.
Slide-39

## Slide 40: Efficiency and thermal Image
Air core
Magnetic core (UCC12050)
Thanks to the 2X peak efficiency, temperature rise of magnetic core solution is ~30°C lower than air-core solution when operating at 5 VIN/5 VOUT100 mA
Notes: We also did comparison of efficiency and thermal performance of aircore solution vs magnetic core solution. The measured peak efficiency of UCC12050 is 2x the air core solution, and a temperature rise of magnetic core solution is ~30°C lower than air-core solution when operating at 5 VIN/5 VOUT at 100 mA.
One important point I want highlight here is about emission. Compared to similar competition product but with aircore, UCC12050 can pass CISPR 32 Class B emissions testing on a two-layer PCB without LDOs or ferrite beads. In an apples-to-apples comparison with competition, UCC12050 has over 20dB of emissions improvement and a margin of close to 5.8 dB.
Slide-40

## Slide 41
ISOWatt – Isolated power and data
Suggested TI device:
ISOW7841
ISOW7821
ISOW7841 schematic
Integrated isolated power & data
Parameter | Value
Input voltage (Vin) | 3V to 5.5V
Output voltage (Vout) | Regulated 3.3V or 5V
Output power (Pout) | 0.65 W
Isolation level | 5kV RMS reinforced
Size | 10.3mm x 7.5mm x 2.65mm
Output regulation | 1%
Frequency
Emission level (dB)
Notes: Another solution to talk about is our isowatt devices. For example, ISOW7841 is reinforced quad-channel digital isolator with integrated DC/DC converter. The digital isolators operate up to 100-Mbps, with a propagation delay less than 16-ns. The integrated DC/DC switch-mode converter supports 130-mA of load current with 5-V input and 75-mA of load current with 3.3-V input.
The emission graph shows comparison of this device with competitor device for CISPR22 Class B.
Slide-41

## Slide 42
Digital isolators – signal isolation
Suggested TI device:
ISO7741DW
ISO7841DWW
ISO7741DW & ISO7841DWW
Application diagram:
Parameter | ISO7741DW | ISO7841DWW
Viso | 5kVrms | 5.7kVrms
Creepage/ Clearance | 8 mm | 14 mm
Data rate | 100 Mbps | 100 Mbps
IEC 60601-1 Capability | 2 MOPP up to 240Vrms | 2 MOPP up to 400Vrms
Size | 10.3 x 7.5 mm | 10.3 x 14.0 mm
SPI Isolation: ISO7741DW vs traditional optocoupler solution
Notes: Before I summarize and compare all the topologies, I wanted to talk about the data or signal isolators. These are mainly designed to replace opto-coupler based isolated designs and I highlight example of SPI isolation in the layout and also the application diagram for the same. The layout shows how much space opto-couplers take and how much space digital isolator takes.
Slide-42

## Slide 43
Layout comparison – power and 4-ch data isolation
Digital isolator (ISO7741) is on bottom layer
1
(ISO7741)
Z = 4.25mm
Z = 6mm
Z = 7.6mm
Notes: Now here is a high-level block diagram comparison of three solutions. First: SN6505 based open-loop push-pull power supply with separate four channel digital isolator. The custom layout takes 30mm x 25mm size.
Second: UCC12050 with separate digital isolator on bottom layer and the overall size is 30mm by 15mm.
Third: ISOW7841 with integrated digital isolators and power possible to layout in same size as UCC1250 i.e. 30mm x 15mm.
There is no specific reason for not including PSR Flyback design. It will take approximately same size as the first option here. Slide-43

## Slide 44
Summary
Topology Parameter | Conventional flyback | PSR flyback (LM25180) | Open-loop push-pull (SN6505) | Closed-loop push-pull (LM25037) | Isolated power module (UCC12050) | Isolated power with digital isolator (ISOW7841)
Output power level | Flexible (transformer and PWM controller dependent) | 5 W to 7 W | 5 W | Flexible (transformer and PWM controller dependent) | 0.5 W | 0.65 W
Input voltage range | Up to 42V/65V | Up to 42V/65V | Up to 5.5V | Up to 75V | Up to 5.5V | Up to 5.5V
Output regulation | 1% or less | 1% | 5 to 10% | 1% or less | 1.5% | 1%
No. of discrete components | More than 30 | 21 | 10 | 46 | Less than 10 | Less than 10
Isolation rating | Flexible (Transformer dependent) | Flexible (Transformer dependent) | Flexible (Transformer dependent) | Flexible (Transformer dependent) | 5000 Vrms Reinforced | 5000 Vrms Reinforced
Emission | High | High | Low | High | Low | Moderate to high
Notes: Finally , here is a comparison summary of all the topologies in terms of different parameters. I don’t want to go through each of them but one thing is important – emissions.
What we have seen is that typically Flyback topologies tend to be noisier than SN6505 (because of Spread Spectrum Clocking). Both conventional & PSR need copper shield for good EMI performance.
UCC12050 uses EMI reduction schemes for internal oscillators and advanced internal layout with proprietary integrated transformer design so EMI is very low.
The EMI performance for ISOWatt is mentioned moderate to high depending on the layout and external components.
All these slides will be shared with everyone on this call however Slide-44

## Slide 45
Application Report
http://www.ti.com/lit/an/sloa285a/sloa285a.pdf
45
Notes: We have also released an application report on TI website which covers all the things in detail. Slide-45

## Slide 46
46
Notes: And finally, if you are not aware, TI has an online E2E community where you can find answers to technical questions, connect with TI factory engineers and also provide your expert comments/suggestions to others.
It is a platform for engineers to connect with engineers.
===
With this I would like to end this presentation. I really appreciate you taking out time to attend this webinar from Texas Instruments. Before we open up for questions, please help to provide your feedbacks on How would you rate this webinar? Was it useful, not so useful?
We can open up for some questions, if there are any.
Slide - 46

## Slide 47
SLYP706
Notes: With this I would like to end this presentation. I really appreciate you taking out time to attend this presentation from Texas Instruments.
We can open up for some questions, if there are any.
Slide - 47