## Slide 1: Design wearable healthcare systems with advanced sensing and efficient power for improved patient monitoring
Sanjay Pithadia
4 March 2021

## Slide 2: Agenda
TI in medical
Market trends
Fundamentals and challenges
Reference designs

## Slide 3
Home healthcare
Medical equipment
Patient monitoring & diagnostics
Medical imaging
TI semiconductors in every medical category
Personal care & fitness
Notes: TI classifies medical market into 5 different categories: Imaging, Patient monitoring and diagnostics, Medical equipment, home healthcare and personal care and fitness.
Medical imaging consists of Ultrasound scanners, X-ray machines, MRI, CT, PET scanners. These are all big machines which has something to do with image of the patient.
Patient monitoring and diagnostics category takes care of all the vital sign sensing and monitoring.
If we consider patient monitoring category as a category which takes human vital signs as “Inputs”, then Medical equipments are mostly “Output devices” like ventilators, infusion pumps, anesthesia machines.
The home healthcare category takes care of devices which are mostly used at home like blood glucose monitors, CPAP machines, nebulizers, oxygen concentrators.
And lastly, Personal healthcare & fitness category has devices for beauty and grooming, wearable fitness monitors, treadmills.
In this presentation, we will talk about the Patient monitoring and diagnostics category.

## Slide 4: Medical sector page

## Slide 5: Medical sensor patches

## Slide 6: Reference design
Comprehensive
design guides
Detailed design
considerations &
applications info
Full
schematics
Design files
TINA-TI
spice simulation
Bill
of
materials
Comprehensive
test results
Boards
available for
evaluation

## Slide 7: Patient monitoring market trend
Remote monitoring enhances quality of care and reduces healthcare cost
Wearable wireless medical technology enables accurate and reliable data in a smaller form factor: multimodalities, longer battery life, SHIP mode
Artificial intelligence uses analytics and big data to improve decision making and early prevention
Notes: The patient monitoring is not new but the market is growing. For example, the growing elderly population drives the high demand of remote monitoring. In Europe, around 700 Billion Euros are being spent every year on chronic diseases. The advantage with remote monitoring enhances the quality of care, reduces overall healthcare cost while providing mobility advantage to the patient.
Wearable wireless medical technology enables accurate and reliable data in a small form factor. The challenges to tackle with wearable medical are: the battery life, the variety of vital signs and modalities and ship mode implementation.
Once the vital signs are monitored, captured and sent for processing. Artificial intelligence can be used for analysis. It can improve the decision making and early prevention of the diseases.

## Slide 8: Patient monitoring basics
The electrocardiogram (ECG)
measures electrical activity of the heart
Photoplethysmography (PPG) is an optical measurement of an organ’s volume.

## Slide 9: ECG vs. PPG
Feature Description | ECG | PPG
Measurement type | Electrical | Optical
Sensor type | Electrodes | Photodiode
Can measure heart rate? | Yes | Yes
Diagnostic information | Yes | Yes
Minimum number of skin contacts required? | 2 (Across chest) | 1 (Finger or wrist)
Number of ADC channels required | ≥1 | 1

## Slide 10: ECG lead and ADC channels
Standards Electrodes Needed
1 Lead LA, RA
3 Lead LA, RA, LL
6 Leads LA, RA, LL
12 Leads LA, RA, LL, V1-6
Number of Leads | Leads Used | Number of ADC Channels
1 | Lead I | 1
3 | Lead I, Lead II, Lead III | 2
6 | Lead I, Lead II, Lead III, aVR, aVL, aVF | 2
12 | Lead I, Lead II, Lead III, aVR, aVL, aVF, V1 – V6 | 8
Notes: A “lead” in ECG is the voltage difference between two points on the body. For most leads, each point is measured by an independent electrode, though some use a calculated average as the second point (but more on this shortly). You can think of different leads as different “viewing angles” that provide unique information about the heart’s electrical activity.
The number of leads in an ECG system can vary widely. Simpler ECG systems generally use no more than the first two or three primary leads. This is typical in consumer applications and small, portable ECG equipment meant for basic diagnostic purposes. On the other hand, high-quality diagnostic systems are more complex and may use up to 15 different leads! The more “viewing angles” included in the ECG, the more diagnostic information the physician can gather about the patient’s heart health.
This table summarizes what leads you measure for single-lead, three-lead, six-lead, and 12-lead ECG, which are the most common configurations. Also shown is how many ADC channels are required to measure each one.
Single-lead ECG is typically measured using Lead I, which uses the two electrodes placed on the arms.
Three lead ECG measures all the limb leads, but only two ADC channels are required since you can calculate the third from the first two.
Six-lead ECG refers to the limb and augmented leads, but still only requires two ADC channels since all the other ECG leads can be calculated from two limb leads.
Finally, 12-lead ECG refers to measuring the limb, augmented and chest leads, but requires only 8 ECG channels since 2 are required for the limb and augmented leads and then 6 are required to measure V1 through V6.

## Slide 11: ECG characteristics Frequency domain

## Slide 12: Challenges in measuring ECG

## Slide 13: Challenges in optical bio-sensing
Low power for longer battery life
Skin tone variation
Best PPG signal for motion cancellation algorithms
Performance with glass
Low temperature performance
Ambient light

## Slide 14: TIDA-01614 Multiparameter front-end for vital signs patient monitor reference design
Benefits
Single IC does both ECG, respiration
Pace detection
ECG with 3 electrodes
Three temperature sensors for temperature measurement
Features
System measures ECG, heart rate, SPO2, respiration rate using ADS1292R and AFE4403 and skin temperature using TMP117
Circuit enables three electrode operation including right leg drive with good CMRR
Pace detection circuit indicates presence of pacemaker
Supports three 0.1º Celsius accurate sensors (TMP117) to measure the skin temperature
Enables data transfer over isolated UART interface
Works with 3.7V Li-ion rechargeable battery
On-board memory for data logging
Applications
Multiparameter patient monitor
Medical sensor patches
Pulse oximeter
Electrocardiogram (ECG)

## Slide 15: Detailed block diagram for TIDA-01614
15

## Slide 16: Design challenges TIDA-01614 solves
Design challenge 2
Protection and isolation
Design challenge 1
Integration of multiple modalities at optimum SNR levels and small form factor
Monitoring of ECG, heart rate, SPO2, PTT, respiration rate and skin temperature
Single lead ECG with RLD (ADS1292R)
> Signal amplitude: 0.2mV~2mV (p-p);
> BW0.05 Hz to 2000 Hz
Supports 3 LED and 3 photodiodes with ambient subtraction for SPO2 and heart rate monitoring with AFE4403
Supports three 0.1º Celsius accurate sensors to measure the skin temperature (TMP117)
Isolated UART interface using an onboard MSP432P401, ISOW7842, TRS3232

## Slide 17: ECG analog front end
  TIDA-01614: Multiparameter front-end reference design for vital signs patient monitor
Important parameters:
Input bias current
Input impedance
Input current noise
Input voltage noise
Power consumption
DC/AC CMRR

## Slide 18: TIDA-01614 test setup and test results
Test setup
GUI display
Design specs

## Slide 19
Test pacemaker detection with TIDA-010005 & TIDA-01614

## Slide 20: TIDA-01580 Wearable, wireless, multiparameter patient monitor reference design
Device Datasheets:
  AFE4900
  CC2640R2F
  TPS61098
  TPS63036
  TPD1E10B06
PPG supports 4 LEDs and 3 PDs with digital ambient subtraction to improve the SNR
AC and DC lead off detection helps in correct measurement of vital signs
Continuous monitoring with lower operating power ensures battery life of 30 days
Flexibility of ultra low power modes and integrated FIFO can keep MCU into sleep to increase the battery operation time
Simple wearable multi-parameter patient monitor for photo-plethysmography (PPG) and electrocardiography (ECG)
Provides raw data to calculate heart-rate, oxygen concentration in blood (SpO2) and pulse-transit time (PTT)
Uses single-chip bio-sensing front-end AFE4900 for synchronized ECG & PPG
  PPG (Optical heart-rate monitoring and SpO2) supports 4 LEDs and 3 PDs with digital ambient subtraction to improve the SNR
  ECG (LEAD I) signals
Integrated ARM Cortex-M3 + 2.4GHz RF transceiver (CC2640R2F) supports wireless data transfer – BLE 4.2 and 5
Operated from CR3032 (3V, 500mA coin cell battery) with battery life of 30 days using highly efficient DC/DC converters
Small form factor helps in easy adaptation to wearable applications
Features
Benefits
Wireless patient monitor
Pulse oximeter
Target applications
Tools & resources
Wearable fitness & activity monitor
ECG

## Slide 21: TIDA-01580 for medical patch
Bottom side is touching the wrist of one hand (ELECTRODE 1)
Other hand is touching the PAD on the top layer of the main board.
(ELECTRODE 2)
Side View
LAUNCHXL-CC2640R2F receives the signals remotely and displays on LabView GUI
The design uses BLE 5.0 with an advertising time = 100ms

## Slide 22: Design challenges TIDA-01580 solves
Design challenge #1
Integration of multiple modalities at optimum SNR levels and small form factor
Capturing synchronized ECG and PPG to enable PTT and BP calculations (non-invasive and without cuff) Pulse transit time (PTT): Time difference between the R-peak in the ECG waveform and the arrival of the blood pressure wave Simultaneous measurement of ECG and PPG together Along with other variables, such as the patient’s size, weight, age, etc., algorithms show the correlation between PTT and systolic blood pressure Challenging to synchronize both measurements – timing is the key! (powering up, clock timing, phase, drift with temperature) | 
22

## Slide 23: Design challenges TIDA-01580 solves
Design challenge # 2
BLE connectivity that does not interfere with measurement accuracy
Signal amplitude: 0.2mV to 2mV (p-p) BW: as broad as 0.05 Hz to 300 Hz (pace detection increases the bandwidth further) Reject environmental electrical signals, such as ac mains, security systems and RFI to amplify and display the ECG signal Good CMRR of the signal chain and right-leg drive (RLD) for CM rejection Differential- and common-mode filtering, environmental shielding and algorithms | 

## Slide 24: Design challenges TIDA-01580 solves
Design challenge # 3
Extending battery life to multiple days to enable portability & wearability
Selecting the right power architecture to enable extended battery life up to 24 hours (for rechargeable batteries) or 7 days (for primary cells) Powering with right buck, boost or buck-boost device instead of directly powering from battery (bypass modes in DC/DC converters) Sleep / shutdown / standby / deep sleep modes for radio devices like BLE, Wi-Fi etc. Selection of right battery charger (charging rate, termination current and quiescent currents are important!) | 
http://www.ti.com/lit/an/slyt763/slyt763.pdf

## Slide 25: TIDA-01624 Bluetooth-enabled high accuracy skin temperature measurement flex PCB patch
TIDA-01624 and/or tools folder
Design guide
Design files: Schematics, BOM, Gerbers, Software, etc.
Device datasheets:
  TMP117
  CC2640R2F
Low power consumption and long battery life
Extremely long shelf life (3+ Years)
Small, flexible form factor
Connects to smart device
Zero-calibration to ±0.1°C accuracy
High accuracy, low power temperature sensor
BLE 4.2 and 5 enabled microcontroller
Thin-film flexible battery power, enabling entirely flexible design
Integrated PCB antenna
Temperature updates every second
Features
Benefits
Tools & resources
Medical sensor patches
Multiparameter patient monitors
Smart patches
Applications
Skin Temperature TMP117
BLE MCU CC2640R2F
I2C

## Slide 26: TMP117x Ultra-high accuracy digital temp sensor with integrated non-volatile memory
Ultra-high accuracy
Meets ASTM E1112 & ISO medical standards:
    0.1°C acc. range 35.8°C to 42°C
No calibration needed; NIST traceable
Integrated non-volatile memory
Store configuration even after losing power
64 bits of general-purpose scratch pad memory
Low power consumption
  3.5uA average Iq @ 1Hz; serial bus inactive
  150nA shutdown Iq; serial bus inactive
Digital feature & I2C interface
Programmable temperature alert & offset value
Soft device rest
Smallest package
6 PIN, QFN & CSP
Features
Benefits
Accuracy
  16-bit resolution (0.0078°C)
  Minimum PSRR: 1LSB = 7.8 m°C/V
Integrated EEPROM
Low power consumption
  140uA Iq during conversion
  3.5uA Average Iq @ 1Hz
  150nA Shutdown Iq
  1.8V – 5.5V
Digital feature: automatic offset NVM/ soft reset
Interface: Single wire
Packaging
  6pin WSON (2 x 2) mm
  6pin WCSP (1.6 x 1) mm
Gas meter
Medical
Cold chain
Wearables
Instrumentation & test
Thermocouple – reference
Applications
TI Part | Accuracy (°C) | Accuracy Full Range
TMP117M | ±0.1oC @ (30oC to 45oC) | ±0.2oC @ (0oC to 85oC)
TMP117 | ±0.1oC @ (-20oC to 50oC) | ±0.3oC @ (-55oC to 150oC)
TMP117N | ±0.2oC @ (-40oC to 100oC) | ±0.3oC @ (-55oC to 150oC)

## Slide 27: Full system: Multiparameter patient monitor + wireless sensors
Patch #1
Patch #2
Patch #3
Patch #4
Nurse’s station / Doctor’s office

## Slide 28: Why TI SimpleLink™ for multiparameter patient monitor + sensor patch?
  BLE SoC with integrated ultra low power sensor controller
  Wi-Fi low power IoT
  Best-in-class standby current
  CC3135/CC3235 Wi-Fi modules
  5GHz Wi-Fi to reliably connect to hospital network
  BLE multirole support, up to 32 simultaneous connections
  FIPS 140-Level 1 validation
  Offload CPU bandwidth – HW crypto accelerators
  Secure boot
  BAW: first crystal-less wireless BLE SoC – 12% area savings in reference design
  Tiny BLE SoC: CC2640R2F – 2.7mm x 2.7mm DSBGA
Secure
Small size
Ease of use
Low power

## Slide 29: Invest once, reuse effortlessly
Learn more about SimpleLink code portability
SimpleLink medical resources
CC2640R2F: How do I design an accurate and thermally efficient wearable temperature monitoring system?
100% code reuse
Common software

## Slide 30: Achieve isolation and help enable patient safety

## Slide 31: Patient safety
Patient safety is a global health priority. Recalling resolution WHA55.18 (2002), which urged Member States to “pay the closest possible attention to the problem of patient safety and to establish and strengthen science-based systems, necessary for improving patients’ safety and the quality of health care”, the seventy-second World Health Assembly (WHA72), in May 2019, adopted WHA72.6, a resolution on ‘Global action on patient safety’. (Source: https://www.who.int/patientsafety/en/)
ECG module
SpO2 module
CO2 module
BP module
Monitor
Notes: The patient monitoring is not new but the market is growing. For example, the growing elderly population drives the high demand of remote monitoring. In Europe, around 700 Billion Euros are being spent every year on chronic diseases. The advantage with remote monitoring enhances the quality of care, reduces overall healthcare cost while providing mobility advantage to the patient.
Wearable wireless medical technology enables accurate and reliable data in a small form factor. The challenges to tackle with wearable medical are: the battery life, the variety of vital signs and modalities and ship mode implementation.
Once the vital signs are monitored, captured and sent for processing. Artificial intelligence can be used for analysis. It can improve the decision making and early prevention of the diseases.

## Slide 32: Isolation requirements and safety limits
IEC60601-1: International basic safety and essential performance standard for electrical medical equipment and medical electrical systems
  Regional compliance
  Editions and versions
Levels of isolation – patient focus
Spacing – creepage & clearances
Safety insulation for transformers
Leakage current limits
  Isolation at the sensing side
  Isolation at the data/power side
Notes: The patient monitoring is not new but the market is growing. For example, the growing elderly population drives the high demand of remote monitoring. In Europe, around 700 Billion Euros are being spent every year on chronic diseases. The advantage with remote monitoring enhances the quality of care, reduces overall healthcare cost while providing mobility advantage to the patient.
Wearable wireless medical technology enables accurate and reliable data in a small form factor. The challenges to tackle with wearable medical are: the battery life, the variety of vital signs and modalities and ship mode implementation.
Once the vital signs are monitored, captured and sent for processing. Artificial intelligence can be used for analysis. It can improve the decision making and early prevention of the diseases.

## Slide 33: Data and power isolation
Notes: The patient monitoring is not new but the market is growing. For example, the growing elderly population drives the high demand of remote monitoring. In Europe, around 700 Billion Euros are being spent every year on chronic diseases. The advantage with remote monitoring enhances the quality of care, reduces overall healthcare cost while providing mobility advantage to the patient.
Wearable wireless medical technology enables accurate and reliable data in a small form factor. The challenges to tackle with wearable medical are: the battery life, the variety of vital signs and modalities and ship mode implementation.
Once the vital signs are monitored, captured and sent for processing. Artificial intelligence can be used for analysis. It can improve the decision making and early prevention of the diseases.

## Slide 34: Key design challenges
Input voltage ranging from 3.3V to 24V
  Regulated input vs. non-regulated input
Output voltage ranging from 3.3V to 6V
Output power up to 5W
Open-loop or closed-loop (voltage/current)
  achieving < 1% load regulation
Isolation ~1kV to 5kV
Emission (CISPR22/25, IEC60601-1)
Small form factor (new trend – electronics in cable and portable MPMs)
  reduced BoM
Low cost
One size doesn’t fit all the requirements
Possible architectures
Flyback
Push-pull
Isolated power module
Isolated power and data module
Notes: The patient monitoring is not new but the market is growing. For example, the growing elderly population drives the high demand of remote monitoring. In Europe, around 700 Billion Euros are being spent every year on chronic diseases. The advantage with remote monitoring enhances the quality of care, reduces overall healthcare cost while providing mobility advantage to the patient.
Wearable wireless medical technology enables accurate and reliable data in a small form factor. The challenges to tackle with wearable medical are: the battery life, the variety of vital signs and modalities and ship mode implementation.
Once the vital signs are monitored, captured and sent for processing. Artificial intelligence can be used for analysis. It can improve the decision making and early prevention of the diseases.

## Slide 35
PSR flyback topology
Tertiary feedback winding needed
Optocoupler needed
TL431
Parameter | Value
Input voltage (Vin) | 4.5V to 65V (70V max)
Output voltage (Vout) | Adjustable
Output power (Pout) | 7W max
Isolation level | 5kV (can be tuned as per transformer design)
Size | 45mm x 25mm x 11mm (Depends on transformer design)
Output regulation | 1% achievable
Suggested TI devices:
LM5180
LM25180
Conventional flyback
Simplified transformer design
No optocouplers required
Smaller size
Total BoM = 21 components (including passives)
Tight load regulation achieved
Primary-side regulated flyback
Refer to “Design Calculator” for complete schematics, BoM and simulation results

## Slide 36: Design calculator for LM25180

## Slide 37
Push-pull topology
Tight output regulation due to feedback
Total BoM = 46 components (including passives)
Optocoupler based design – reliability
Push-pull topology (closed-loop)
Push-pull topology (open-loop)
PMP design: PMP40588
5-10% regulation
No opto-couplers required
Smaller size, total BoM = 10 components (including passives)
Needs regulated input
Parameter | Value
Input voltage (Vin) | 2.2V to 5.5V
Output voltage (Vout) | 5V unregulated
Output power (Pout) | 5W max
Isolation level | 5kV (can be tuned as per transformer design)
Size | 30mm x 25mm x 6mm (Depends on transformer design)
Output regulation | 5 to 10%
Suggested TI devices:
SN6505A
SN6505B

## Slide 38
Isolated power module
Vin
Vout
Suggested TI device:
UCC12050
Same (apple-to-apple) EVM configuration: no ferrite beads , no LDO, no stitch capacitors, on 2 layer PCB
Tested to CISPR32 Limit, in 10m chamber, on same day, in same certified lab.
UCC12050 schematic
Scalability
Parameter | Value
Input voltage (Vin) | 4.5V to 5.5V
Output voltage (Vout) | Regulated 3.3V or 5V
Output power (Pout) | 0.5 W
Isolation level | 5kV RMS reinforced
Size | 10.3mm x 7.5mm x 2.65mm
Output regulation | 1.5%

## Slide 39
Integrated transformer technology benefits
Discrete solution
38mm
21mm
7.5mm
Total bias area
273 mm^2
53%
less area
Total bias area
127 mm^2
UCC12050
3.5mm
Single chip solution (UCC12050) advantages:
Smaller size and low profile
Very low isolation capacitance Cps for better CMTI and less noise
Simplify design with less components and easy board layout
Digital isolator
Isolated bias UCC12050

## Slide 40: Efficiency and thermal Image
Air core
Magnetic core (UCC12050)
Thanks to the 2X peak efficiency, temperature rise of magnetic core solution is ~30°C lower than air-core solution when operating at 5 VIN/5 VOUT100 mA

## Slide 41
ISOWatt – Isolated power and data
Suggested TI device:
ISOW7841
ISOW7821
ISOW7841 schematic
Integrated isolated power & data
Parameter | Value
Input voltage (Vin) | 3V to 5.5V
Output voltage (Vout) | Regulated 3.3V or 5V
Output power (Pout) | 0.65 W
Isolation level | 5kV RMS reinforced
Size | 10.3mm x 7.5mm x 2.65mm
Output regulation | 1%
Frequency
Emission level (dB)

## Slide 42
Digital isolators – signal isolation
Suggested TI device:
ISO7741DW
ISO7841DWW
ISO7741DW & ISO7841DWW
Application diagram:
Parameter | ISO7741DW | ISO7841DWW
Viso | 5kVrms | 5.7kVrms
Creepage/ Clearance | 8 mm | 14 mm
Data rate | 100 Mbps | 100 Mbps
IEC 60601-1 Capability | 2 MOPP up to 240Vrms | 2 MOPP up to 400Vrms
Size | 10.3 x 7.5 mm | 10.3 x 14.0 mm
SPI Isolation: ISO7741DW vs traditional optocoupler solution

## Slide 43
Layout comparison – power and 4-ch data isolation
Digital isolator (ISO7741) is on bottom layer
1
(ISO7741)
Z = 4.25mm
Z = 6mm
Z = 7.6mm

## Slide 44
Summary
Topology Parameter | Conventional flyback | PSR flyback (LM25180) | Open-loop push-pull (SN6505) | Closed-loop push-pull (LM25037) | Isolated power module (UCC12050) | Isolated power with digital isolator (ISOW7841)
Output power level | Flexible (transformer and PWM controller dependent) | 5 W to 7 W | 5 W | Flexible (transformer and PWM controller dependent) | 0.5 W | 0.65 W
Input voltage range | Up to 42V/65V | Up to 42V/65V | Up to 5.5V | Up to 75V | Up to 5.5V | Up to 5.5V
Output regulation | 1% or less | 1% | 5 to 10% | 1% or less | 1.5% | 1%
No. of discrete components | More than 30 | 21 | 10 | 46 | Less than 10 | Less than 10
Isolation rating | Flexible (transformer dependent) | Flexible (transformer dependent) | Flexible (transformer dependent) | Flexible (transformer dependent) | 5000 Vrms reinforced | 5000 Vrms reinforced
Emission | High | High | Low | High | Low | Moderate to high

## Slide 45
Application report
http://www.ti.com/lit/an/sloa285a/sloa285a.pdf
45

## Slide 46
46
Notes: Lastly, you can go and visit e2e.ti.com which is TI engineer to engineer community for all the technical queries. It has many questions answered already.
We hope that you enjoyed this webinar. Now we open up for the Q&A. You can send your questions to Sébastien Launay.

## Slide 47
©2021 Texas Instruments Incorporated. All rights reserved.
SLYP747